## Slide 1: OL 750 Measurement Systems

## Slide 2: OUTLINE OF PRESENTATION • About Optronic Laboratories, LLC • The OL 750 Series Spectroradiometer • Source Measurements • Transmission Measurements • Reflectance Measurements • Detector Measurements

## Slide 3: ABOUT OPTRONIC LABORATORIES, LLC • Optronic Laboratories, LLC was established in 1970 by two eminent researchers at NIST (then NBS). Facilities were modeled after NIST, and Optronic Laboratories, LLC achieved a worldwide reputation for excellence in light measurement and calibration. • Continuing to work closely with NIST, the development of new instruments and techniques at Optronic Laboratories, LLC constantly improves the accuracy and precision of light measurement.
BACKGROUND

## Slide 4: OPTRONIC LABORATORIES’ INSTRUMENTS & SERVICES INCLUDE: • Training, consultation & advice
BACKGROUND

## Slide 5: THE HEART OF ANY SPECTRORADIOMETRIC SYSTEM IS THE MONOCHROMATOR, BUT… • It must give the right bypass – - Too large and spectra are distorted - Too small and signals are low • It must give high spectral purity (low stray light) so only the wavelength of interest is detected
THE OL SERIES 750 MONOCHROMATORS

## Slide 6: • Available in single and double (shown) versions • F/4 optics for high throughput • Very low stray light • Numerous special features for most accurate results
THE OL SERIES 750 MONOCHROMATORS

## Slide 7: • Large range of gratings available • Optimized for high efficiency • Optimized for wide wavelength range • Optimized for high resolution
THE OL SERIES 750 MONOCHROMATORS
Tri-grating mounts change gratings automatically

## Slide 8: • Range of plug-in input accessories • Choppers (shown) for IR detectors • Automatic control of chopper speed and “open” position stop • Filter-wheel for input attenuation or filtering
THE OL SERIES 750 MONOCHROMATORS
Versatile plug-in chopper or filter wheel input accessories

## Slide 9: • Fully automated 11 position blocking-filter wheel • Includes SHUT position for dark current readings • Optimized filters for widest wavelength range
THE OL SERIES 750 MONOCHROMATORS
Blocking-filter wheel changes filters automatically

## Slide 10: • Gratings diffract light at different angles depending on wavelength • 400nm will also diffract at the same angle as 800nm (since 2 x 400 = 800) • A blocking filter removes the 400nm so just 800nm is left
THE OL SERIES 750 MONOCHROMATORS
Blocking filters:
how they work
400nm + 800nm
400nm
400nm
800nm
Blocking filter
grating

## Slide 11: • Light passes into the monochromator through the entrance slit, then onto the first collimating mirror.
THE OL SERIES 750 MONOCHROMATORS

## Slide 12: • Collimated (parallel) light is directed onto the grating.
THE OL SERIES 750 MONOCHROMATORS

## Slide 13: • The grating diffracts the light, sending each wavelength off at a different angle. • The desired wavelength is collected by the first focussing mirror….
THE OL SERIES 750 MONOCHROMATORS

## Slide 14: • …which focuses it at the middle slit. • A second collimating mirror collects this beam…
THE OL SERIES 750 MONOCHROMATORS

## Slide 15: • …and collimates it onto a second grating. • Since the second grating is linked to the first, they are at the same angle and hence the same wavelength.
THE OL SERIES 750 MONOCHROMATORS

## Slide 16: • This second grating acts like the first, but the spectral purity is increased by many decades.
THE OL SERIES 750 MONOCHROMATORS

## Slide 17: • The pure monochromatic light, at the wavelength desired, is then focussed onto the exit slit and emerges from the monochromator
THE OL SERIES 750 MONOCHROMATORS

## Slide 18: THE MOST COMMON MEASUREMENTS OF SOURCES INCLUDE: • Radiance - Light emitted in a particular direction • Irradiance – Light falling onto a surface • Total flux – Total light emitted in all directions
SOURCE MEASUREMENT

## Slide 19: • Radiance is normally measured with a telescope • The solid collection angle corresponds to the size of the lens • The measurement area corresponds to the field-of- view of the telescope
RADIANCE MEASUREMENT
Solid Collection angle
Measurement area
Source

## Slide 20: • The term “radiance” in this section also refers to related units such as luminance and radiant intensity • Radiant intensity applies to small or distant objects that are smaller than the telescope field-of-view
RADIANCE MEASUREMENT
Source
Measurement area
Solid Collection angle

## Slide 21: • Optronic Laboratories, Inc. manufactures a large range of telescope, microscope and imaging optic modules for radiance measurements • To illustrate a typical system for radiance measurement, we will use the OL Series 730-9 Telescopes
RADIANCE MEASUREMENT

## Slide 22: • For luminance and radiance of sources from 200 nm to 14,000 nm
THE OL SERIES
730-9 TELESCOPES
Wide range of lenses
Reflex viewing
Adjustable focus

## Slide 23: • For luminance and radiance of sources from 200 nm to 14,000 nm
THE OL SERIES
730-9 TELESCOPES
Light coming from the source…
Here is how it works

## Slide 24: • For luminance and radiance of sources from 200 nm to 14,000 nm
THE OL SERIES
730-9 TELESCOPES
…is picked up at some distance and focused by the OL 730-9
The sectional drawing shows what happens inside the solid housing
If the mirror is flipped out of the way

## Slide 25: • For luminance and radiance of sources from 200 nm to 14,000 nm
THE OL SERIES
730-9 TELESCOPES
The image is sent via the field-of-view aperture to the spectroradiometer

## Slide 26: • Putting together an OL 750 radiance system
RADIANCE MEASUREMENT
Then we add the OL 730-9
And a detector
Positive location & easy assembly
We start with an OL 750 monochromator

## Slide 27
RADIANCE MEASUREMENT
We have a complete radiance measurement system

## Slide 28: • To measure irradiance (or illuminance), a cosine response integrating sphere such as the OL IS-670 is substituted for the OL 730-9 • To measure total flux, a large sphere (typically 18 to 76 inches) is substituted for the OL 730-9
IRRADIANCE AND
TOTAL FLUX MEASUREMENT

## Slide 29
IRRADIANCE MEASUREMENT
A complete irradiance measurement system
OL IS-670 near-perfect cosine response integrating sphere

## Slide 30: • For transmission, reflection and detector response measurements, a light source at the entrance of the OL 750 creates monochromatic light at the exit • The OL Series 740-20 is an ideal light source for most applications
OTHER MEASUREMENTS

## Slide 31: • Single and double lamp versions • UV lamp - 200nm to 400nm • Tungsten lamp - 250nm to 2500nm • Glowbar - 1000nm to 30000nm
THE OL SERIES 740-20
LIGHT SOURCES

## Slide 32: • Optics match the OL Series 750 Monochromators • Easy switching from one source …
THE OL SERIES 740-20
LIGHT SOURCES

## Slide 33: • Optics match the OL Series 750 Monochromators • Easy switching from one source … …TO THE OTHER
THE OL SERIES 740-20
LIGHT SOURCES
Fully automated switching available

## Slide 34: • 200 to 30,000 nm coverage • Monochromatic light prevents sample heating • Suits a very wide range of standard OL accessories • Suits many existing and future applications
THE OL 750 + OL 740-20

## Slide 35: • Transmittance can be regular (like a window) or diffuse (like paper) • Optronic Laboratories, Inc. manufacture a wide range of accessories for regular and diffuse transmittance • To illustrate a regular transmittance measurement system, we will use the OL Series 740-73 Transmittance Attachment
TRANSMITTANCE
MEASUREMENT

## Slide 36: • After a 100% scan (with setup as shown), the sample is placed in the beam
THE OL SERIES 740-73
Collimated light
Detector
Sample

## Slide 37: • The ratio of a scan with the sample in place to the 100% scan is the transmittance of the sample
THE OL SERIES 740-73

## Slide 38: • When the system is assembled, the transmission of a sample can be measured up to 14,000 nm
THE OL 740-73 +
OL 750 + OL 740-20
Other transmittance accessories are available to cover the range
200 nm to 30,000nm

## Slide 39: • Reflectance can be specular (like a mirror) or diffuse (like paper) • Optronic Laboratories, Inc. manufacture a wide range of accessories for regular and diffuse transmittance • To illustrate a specular reflectance measurement system, we will use the OL 750-75MA Goniometer Attachment
REFLECTANCE
MEASUREMENT

## Slide 40: • Measures direct beam for true absolute results without the need for standard mirrors
THE OL 750-75MA
FOR SPECULAR REFLECTANCE

## Slide 41: • Full 10° to 80° angle-of- incidence measurements • Detector tracks incidence angle or moves free to map other angles
THE OL 750-75MA
FOR SPECULAR REFLECTANCE

## Slide 42: • Transmittance at any angle, 0° to 80°
THE OL 750-75MA FOR SPECULAR TRANSMITTANCE

## Slide 43: • Detector moves independent of incident angle for BRDF & BTDF measurements
THE OL 750-75MA
FOR BRDF MEASUREMENTS

## Slide 44: • If we put these together…
THE OL 750 + OL 740-20
+ OL 750-75MA
Positive location
&
Easy assembly

## Slide 45: • Materials can be measured from 200nm to 30,000nm
THE OL 750 + OL 740-20
+ OL 750-75MA

## Slide 46: • The OL 750-75MA can be used to map diffuse reflectance by measuring one angle at a time • To give diffuse reflectance in one measurement however, an integrating sphere is required
DIFFUSE
REFLECTANCE

## Slide 47: • Optronic Laboratories, LLC manufacture a range of fixed-angle and Edward Spheres for diffuse reflectance measurements. • To illustrate a diffuse reflectance measurement system, we will use the OL 750-70 Diffuse Reflectance Attachment.
DIFFUSE
REFLECTANCE

## Slide 48: • Diffuse reflectance is measured against a standard • A scan is run with the standard, then another with the sample
DIFFUSE
REFLECTANCE
Sample

## Slide 49: • From a simple viewpoint, the ratio of these scans would indicate the sample reflectance • However, the throughput of the sphere changes due to the sample
DIFFUSE
REFLECTANCE
Sample
CAUTION:
Change in throughput
DO NOT USE

## Slide 50: • Instead, the double beam method should be used • This involves 4 scans, 3 to calibrate and one to measure
DIFFUSE
REFLECTANCE
Sample
Scan 3
Scan 4

## Slide 51: • The scans with the beam directed at the wall of the sphere correct for the throughput changes • The true reflectance can then be calculated
DIFFUSE
REFLECTANCE
Sample
Scan 4

## Slide 52: • Uses the double beam method for accurate values of reflectance • Software guides the user through the calibration and measurement
THE OL SERIES
750-70
Standard or sample
Beam switch mirror

## Slide 53: • 2 versions of the OL 750-70 are available - PTFE sphere coating for 250nm to 2,500nm - Gold sphere coating for 1,000nm to 30,000nm
THE OL SERIES
750-70
Standard or sample
Beam switch mirror

## Slide 54: • If we put these together…
THE OL 750 + OL 740-20
+ OL 750-70
And add a detector
Positive location
&
Easy assembly

## Slide 55: • We have a complete system for diffuse reflectance
THE OL 750 + OL 740-20
+ OL 750-70
The light-trap supplied also allows measurements which exclude specular reflections

## Slide 56: • In order to characterize their performance in a particular applications, single detectors, array detectors and assemblies often need spectral response to be known • For example, photopic detectors can be analysed for V() matching
DETECTOR
SPECTRAL RESPONSE

## Slide 57: • Measurements are generally consistent with detector use, and may require a large collimated beam, a small spot, a diffuse source or other geometry • Optronic Laboratories, LLC manufactures a wide range of accessories for most detector measurement requirements, including quantum efficiency
DETECTOR
SPECTRAL RESPONSE

## Slide 58: • Optronic Laboratories, LLC provides standard detectors, traceable to NIST and other National Laboratories • To illustrate a detector spectral response system, we will use the OL 750-10C All Mirror Collimator
DETECTOR
SPECTRAL RESPONSE

## Slide 59: • 2 inch diameter beam • Excellent beam uniformity • Constant collimation at all wavelengths
THE OL 750-10C

## Slide 60: • If we put these together…
THE OL 750 + OL 740-20
+ OL 750-10C

## Slide 61: • We have a complete system for detector response
THE OL 750 + OL 740-20
+ OL 750-10C
Then we measure the test detector

## Slide 62: • The system illustrated represents just a few of the accessories from Optronic Laboratories, LLC • For more information: Visit our website at OptronicLabs.com Call 407-422-3171 to Contact an Application Engineer
MORE INFORMATION