## Slide 1: OL 750 Diffuse Spectral Reflectance Measurement System

## Slide 2: ABOUT OPTRONIC LABORATORIES, LLC • In 1998, Optronic Laboratories, LLC moved into its new 25,000 square foot purpose-built facility, consolidating and expanding production, R & D and calibration facilities.
BACKGROUND

## Slide 3: THE HEART OF ANY SPECTRORADIOMETRIC SYSTEM IS THE MONOCHROMATOR, BUT… • It must give the right bypass – - Too large and spectra are distorted - Too small and signals are low • It must give high spectral purity (low stray light) so only the wavelength of interest is detected
THE OL SERIES 750 MONOCHROMATORS

## Slide 4: • Available in single and double (shown) versions • F/4 optics for high throughput • Very low stray light • Numerous special features for most accurate results
THE OL SERIES 750 MONOCHROMATORS

## Slide 5: • Large range of gratings available • Optimized for high efficiency • Optimized for wide wavelength range • Optimized for high resolution
THE OL SERIES 750 MONOCHROMATORS
Tri-grating mounts change gratings automatically

## Slide 6: • Range of plug-in input accessories • Choppers (shown) for IR detectors • Automatic control of chopper speed and “open” position stop • Filter-wheel for input attenuation or filtering
THE OL SERIES 750 MONOCHROMATORS
Versatile plug-in chopper or filter wheel input accessories

## Slide 7: • Fully automated 11 position blocking-filter wheel • Includes SHUT position for dark current readings • Optimized filters for widest wavelength range
THE OL SERIES 750 MONOCHROMATORS
Blocking-filter wheel changes filters automatically

## Slide 8: • Gratings diffract light at different angles depending on wavelength • 400nm will also diffract at the same angle as 800nm (since 2 x 400 = 800) • A blocking filter removes the 400nm so just 800nm is left
THE OL SERIES 750 MONOCHROMATORS
Blocking filters:
how they work
400nm + 800nm
400nm
400nm
800nm
Blocking filter
grating

## Slide 9: • Light passes into the monochromator through the entrance slit, then onto the first collimating mirror.
THE OL SERIES 750 MONOCHROMATORS

## Slide 10: • Collimated (parallel) light is directed onto the grating.
THE OL SERIES 750 MONOCHROMATORS

## Slide 11: • The grating diffracts the light, sending each wavelength off at a different angle. • The desired wavelength is collected by the first focussing mirror….
THE OL SERIES 750 MONOCHROMATORS

## Slide 12: • …which focuses it at the middle slit. • A second collimating mirror collects this beam…
THE OL SERIES 750 MONOCHROMATORS

## Slide 13: • …and collimates it onto a second grating. • Since the second grating is linked to the first, they are at the same angle and hence the same wavelength.
THE OL SERIES 750 MONOCHROMATORS

## Slide 14: • This second grating acts like the first, but the spectral purity is increased by many decades.
THE OL SERIES 750 MONOCHROMATORS

## Slide 15: • The pure monochromatic light, at the wavelength desired, is then focussed onto the exit slit and emerges from the monochromator
THE OL SERIES 750 MONOCHROMATORS

## Slide 16: • For transmission, reflection and detector response measurements, a light source at the entrance of the OL 750 creates monochromatic light at the exit. • The OL Series 740-20 is an ideal light source for most applications.
OL SERIES 740-20
LIGHT SOURCES

## Slide 17: • Single and double lamp versions • UV lamp - 200nm to 400nm • Tungsten lamp - 250nm to 2500nm • Glowbar - 1000nm to 30000nm
THE OL SERIES 740-20
LIGHT SOURCES

## Slide 18: • Optics match the OL Series 750 Monochromators • Easy switching from one source …
THE OL SERIES 740-20
LIGHT SOURCES

## Slide 19: • Optics match the OL Series 750 Monochromators • Easy switching from one source … …TO THE OTHER
THE OL SERIES 740-20
LIGHT SOURCES
Fully automated switching available

## Slide 20: • 200 to 30,000 nm coverage • Monochromatic light prevents sample heating • Suits a very wide range of standard OL accessories • Suits many existing and future applications
THE OL 750 + OL 740-20

## Slide 21: • Reflectance can be specular (like a mirror) or diffuse (like paper) • Optronic Laboratories, Inc. manufacture a wide range of accessories for regular and diffuse transmittance
REFLECTANCE
MEASUREMENT

## Slide 22: • Optronic Laboratories, Inc. manufacture a range of fixed-angle and Edward Spheres for diffuse reflectance measurements. • To illustrate a diffuse reflectance measurement system, we will use the OL 750-70 Diffuse Reflectance Attachment.
DIFFUSE
REFLECTANCE

## Slide 23: • Diffuse reflectance is measured against a standard • A scan is run with the standard, then another with the sample
DIFFUSE
REFLECTANCE
Sample

## Slide 24: • From a simple viewpoint, the ratio of these scans would indicate the sample reflectance • However, the throughput of the sphere changes due to the sample
DIFFUSE
REFLECTANCE
Sample
CAUTION:
Change in throughput
DO NOT USE

## Slide 25: • Instead, the double beam method should be used • This involves 4 scans, 3 to calibrate and one to measure
DIFFUSE
REFLECTANCE
Sample
Scan 3
Scan 4

## Slide 26: • The scans with the beam directed at the wall of the sphere correct for the throughput changes • The true reflectance can then be calculated
DIFFUSE
REFLECTANCE
Sample
Scan 4

## Slide 27: • Uses the double beam method for accurate values of reflectance • Software guides the user through the calibration and measurement
THE OL SERIES
750-70
Standard or sample
Beam switch mirror

## Slide 28: • 2 versions of the OL 750-70 are available - PTFE sphere coating for 250nm to 2,500nm - Gold sphere coating for 1,000nm to 30,000nm
THE OL SERIES
750-70
Standard or sample

## Slide 29: • If we put these together…
THE OL 750 + OL 740-20
+ OL 750-70
And add a detector
Positive location
&
Easy assembly

## Slide 30: • We have a complete system for diffuse reflectance
THE OL 750 + OL 740-20
+ OL 750-70
The light-trap supplied also allows measurements which exclude specular reflections

## Slide 31: • The system illustrated represents just a few of the accessories from Optronic Laboratories, LLC • For more information: Visit our website at OptronicLabs.com Call 407-422-3171 to Contact an Application Engineer
MORE INFORMATION